## Slide 1
Deep Learning-Based PET Image Correction Toward Quantitative Imaging
Presenter: Sama Shahpouri
Study : Data Science for Life Science
Institute: Hanze University of Applied Sciences, Institute of Life Science & Technology
Supervisors: Dr. Isaac Shiri, Dr. Dave Landers, Dr. Gijs van der Schot
28/08/2024

## Slide 2: Introduction of PET & ASC Problem Statement Research Objective Data Collection Methodology Internal Validation Results External Validation Results Artifact Correction Results Conclusion
Agenda

## Slide 3
https://youtu.be/oc3HRxemNGI?si=djrXXO1kcawsbd8_

## Slide 4
PET Attenuation and Scatter Correction (ASC)
Attenuation correction (AC) of annihilation photons along lines of response (LOR) in the patient body is another critical step
Scatter correction (SC) of Compton scattered photons, which can result in mis-allocation of the annihilated photons, is crucial
True Coincidence
One annihilation
Straight path photons
Scatter Coincidence
One annihilation
Photons scatter
Artifactual measured LOR
Random Coincidence
More than one annihilation
Photons from different annihilations are detected simultaneously
Artifactual measured LOR
https://www.radiologycafe.com/frcr-physics-notes/molecular-imaging/pet-imaging/

## Slide 5
The non-attenuation corrected image (NAC) PET: with ASC (accurate) CT image (accur
Enhanced skin uptake
reduced mediastinal
uptake
Non-uniform liver
'hot' lungs
Effects of Attenuation Scatter Correction

## Slide 6
Data flow in a PET/CT scanner
Schmitz et al. The physics of PET/Ct scanners. PET and PET/CT: a clinical guide, 2005.
X-ray acquisition
Anatomical (CT) Reconstruction
PET Emission Acquisition
CT Image
Translate CT to PET Energy (511 keV)
Smooth to PET Resolution
Attenuation/Scatter Correct PET Emission Data
Functional (PET) Reconstruction
PET Image

## Slide 7: Challenges in ASC using CT info
☢️Radiation Exposure
⚠️CT Artifacts Impact PET Data
🔄Misalignment Issues
🛠️Metal Artifacts
🌬️Respiratory Pattern Differences
🚫Limited Availability
💰Cost and Complexity
💉Contrast Agent Issues
✂️CT Truncation Effects
🔦CT Bias Due to Beam-Hardening

## Slide 8
68Ga-labelled pharmaceuticals present challenges due to low tracer activity and extensive positron range.

## Slide 9
Types of common Artifacts
Metallic
Artifact
Truncation Artifacts
Motion Artifacts
Hallo Artifacts

## Slide 10
AI-based ASC in PET
Indirect approaches use MRI or CT to generate attenuation maps.
GANs have been used to create pseudo-CT images from NAC PET scans.
DL-enhanced MLAA algorithm reduces common issues like crosstalk artifacts and noisy attenuation maps.
Direct DL-ASC methods generate ASC PET images directly from NAC images.

## Slide 11: ⇓Research Objective
✔️ Address direct ASC of PET images without CT anatomical data
✔️ Create a universal model applicable to different radiotracers and scanners ✔️ Evaluate model performance in artifact removal across a multi-center dataset

## Slide 12: Materials and Methods

## Slide 13: Workflow of A CT-less deep convolutional neural network
PET raw data
CT images
Recon without AC & SC
NAC (Input)
Model Training
DL PET
Recon with AC & SC
MAC (Target)
Recon = reconstruction;
AC = attenuation correction,
SC = scatter correction,
NAC = noncorrected PET,
MAC = CT-Measured attenuation/scatter corrected,
DL PET = deep learning–corrected PET,

## Slide 14: Workflow of Decomposition-Based Approach [1]
PET raw data
CT images
Recon without AC & SC
Recon with AC & SC
NAC (Input)
Model Training
DL PET
MAC (Target)
[2] Guo R, Xue S, Hu J, Sari H, Mingels C, Zeimpekis K, et al. Using domain knowledge for robust and generalizable deep
learning-based CT-free PET attenuation and scatter correction. Nat Commun. 2022 Dec 1;13(1).

## Slide 15
[2] Guo R, Xue S, Hu J, Sari H, Mingels C, Zeimpekis K, et al. Using domain knowledge for robust and generalizable deep
learning-based CT-free PET attenuation and scatter correction. Nat Commun. 2022 Dec 1;13(1).
NAC (Input)
Model Training
DL PET
MAC (Target)
Anatomy-Dependent Correction Maps (ADCM)
Anatomy-independent textures
Workflow of Decomposition-Based Approach [1]

## Slide 16: Workflow of Decomposition-Based Approach [1]
[2] Guo R, Xue S, Hu J, Sari H, Mingels C, Zeimpekis K, et al. Using domain knowledge for robust and generalizable deep
learning-based CT-free PET attenuation and scatter correction. Nat Commun. 2022 Dec 1;13(1).
NAC (Input)
ADCM
Model Training
DL PET
Anatomy-Dependent Correction Maps (ADCM)
Anatomy-independent textures

## Slide 17: Workflow of Decomposition-Based Approach [1]
[2] Guo R, Xue S, Hu J, Sari H, Mingels C, Zeimpekis K, et al. Using domain knowledge for robust and generalizable deep
learning-based CT-free PET attenuation and scatter correction. Nat Commun. 2022 Dec 1;13(1).
NAC (Input)
ADCM
Model Training
DL ADCM
DL PET

## Slide 19
68Ga PET/CT Dataset
Shiri I, Salimi Y, Hervier E, Pezzoni A, Sanaat A, Mostafaei S, et al. Artificial Intelligence-Driven Single-Shot PET Image Artifact Detection and Disentanglement: Toward Routine Clinical Image Quality Assurance. Clin Nucl Med. 2023 Dec 1;48(12):1035–46
Center | No | Train | Validation | Test | Scanner | Reconstruction | Matrix size × Z*
Center 1 | 56 | 43 | 11 | 2 | Siemens Biograph 6 | 3D-OSEM | 168 × 168
Center 2 | 31 | 25 | 4 | 2 | GE Discovery IQ | 3D-OSEM | 192 × 192
Center 3 | 45 | 35 | 8 | 2 | Siemens mCT | 3D-OSEM | 200 × 200
Center 4 | 40 | 28 | 10 | 2 | Siemens Biograph 6 | 3D-OSEM | 168 × 168
External Center | 12 | - | - | 12 | Siemens Horizon | PSF+TOF+3D-OSEM | 180 × 180
Total | 184 | 131 | 33 | 20 | - | - | -

## Slide 20
18F-FDG Dataset
Center | No | Train | Validation | Test | Scanner | Reconstruction | Matrix size × Z*
Center 6 | 55 | 39 | 6 | 11 | Siemens Biograph 6 | 3D-OSEM | 168 × 168
Center 7 | 43 | 23 | 9 | 10 | GE Discovery IQ | 3D-OSEM | 192 × 192
Total | 98 | 62 | 15 | 21 | - | - | -

## Slide 21: 198 patients 18Ga PSMA PET/CT Motion and Halo artifacts
Artifact dataset

## Slide 22: Test Sets
Internal Center: 10% of 4 centers of Ga dataset
External Center: one of the 5 centers of Ga dataset
External Tracer (FDG): 10% of 2 centers FDG dataset
Artifact images: All of the dataset

## Slide 23: Hyper parameters
Dyn-UNet architecture
3D patches sized 168×168×16
20 patches per patient
Loss function of the mean squared error (MSE)
Learning rate of 0.001
Optimization Adam algorithm
Batch size: 4
spacing : [4.07, 4.07, 3.00]
four layers: [16, 32, 64, 128]
Kernels: [3, 3, 3]
2 super vision head
The architecture of DynUnet

## Slide 24
Integrated multi-Centre model (IMCM): the best metric of 0.0527 at epoch 434.
18F-FDG

## Slide 25
Anatomy-Dependent Correction Model (ADCM): the best metric of 0.1237 was reached at epoch 466.
68Ga-PET

## Slide 26
Transfer Learning for IMCM model (TL-MC):
18F-FDG

## Slide 27
18F-FDG

## Slide 28: Results

## Slide 29
68Ga-PET
Cross-Center

## Slide 30
68Ga-PET
Cross-Center

## Slide 31
18F-FDG
Cross-Tracer

## Slide 32
18F-FDG
Cross-Tracer

## Slide 33: Case Study on artifact images

## Slide 35
Coronal views of 6 clinical studies, representing from left to right: NAC, MAC, IMCM-DL and the difference images of MAC and DL image. Our method effectively disentangles halo artefacts in the kidney area.

## Slide 36
Coronal views, showing from left to right NAC, MAC, IMCM-DL and the difference images of MAC and DL image. The images generated using the IMCM approach successfully corrected the mismatch artefact in the diaphragm, lung, liver and spleen regions.

## Slide 39
Developed CT-free DL models (IMCM and ADCM) for ASC in PET imaging, tested across different centers and radiotracers.
Single universal models may not be effective; radioisotope-specific tuning is necessary.
Using large, heterogeneous datasets enhances ASC accuracy in PET images.
IMCM showed significant improvement in quantitative accuracy and artifact correction.
IMCM outperformed ADCM, demonstrated more reliable and consistent performance in SUV estimations.
Discussion

## Slide 40
Conclusion

## Slide 41
DL methods improve both quantitative and qualitative aspects of PET images, detecting and correcting artifacts without anatomical images.
Designing a neural network for applying ASC on PET images without using CT data.
Working well in different centers, making it reliable and flexible for use in clinics.
challenging the model's capability in detecting and removing artifacts in PET scans.
Created special models for each site using transfer learning.

## Slide 42
Limitations

## Slide 43
Future Work
Future research should focus on clinical imaging parameters and organ-specific evaluations.
Different neural networks and optimizers could improve model performance.
Model performance must be evaluated across various centers for generalization and robustness.
Clinical Imaging Parameters:
Explore SUVmean, SUVmax, and total lesion metabolism.
Assess radiomic features within the sphere of influence.
Organ-Specific Evaluations:
Targeted analysis for clean and artifactual images.
Rigorous statistical testing of categorized outcomes.
Alternative Neural Networks:
Investigate generative adversarial networks (GANs), variational autoencoders, transformer-based models, or Swin-UNet.
Hyperparameter Optimization:
Experiment with optimizers like AdamW, RMSprop, and Nadam.
Use tools like Optuna for systematic hyperparameter tuning.
Multi-Center Evaluation:
Test model performance on data from different centers.
Ensure consistency and reliability across various clinical environments.