## Slide 1
Deep Learning-Based PET Image Correction Toward Quantitative Imaging
Presenter: Sama Shahpouri
Study : Data Science for Life Science
Institute: Hanze University of Applied Sciences, Institute of Life Science & Technology
Supervisors: Dr. Isaac Shiri, Dr. Dave Landers, Dr. Gijs van derhot
28/08/2024

## Slide 2: Introduction

## Slide 3: PET Attenuation and Scatter Correction (ASC)
Attenuation correction (AC) of annihilation photons along lines of response (LOR) in the patient body is another critical step
Scatter correction (SC) of Compton scattered photons, which can result in mis-allocation of the annihilated photons, is crucial
Notes: True Coincidence
One annihilation
Straight path photons
Scatter Coincidence
One annihilation
Photons scatter
Artifactual measured LOR

## Slide 4: Effects of Attenuation Scatter Correction
Enhanced skin uptake
Non-uniform liver
'Hot' lungs
Reduced mediastinal
uptake
Non-Attenuation Corrected image (NAC) Attenuation-scatter corrected image(MAC) CT image
Notes: MAC (Model-based Attenuation and Scatter Correction)

## Slide 5: Data flow in a PET/CT scanner [1]
X-ray acquisition
Anatomical (CT) Reconstruction
PET Emission Acquisition
CT Image
Translate CT to PET Energy (511 keV)
Smooth to PET Resolution
Attenuation/Scatter Correct PET Emission Data
Functional (PET) Reconstruction
PET Image

## Slide 6: Challenges in ASC using CT info
☢️Radiation Exposure
🚫Limited Availability
💰Cost and Complexity
💉Contrast Agent Issues
🩻 Artifacts

## Slide 7: Challenges in ASC using CT info
Halo Motion Truncation Metal Mismatch
[4] Shiri, Clin Nucl Med. 2023

## Slide 8: AI is starting to change medical imaging, for real
McMillan, PET Clinic, 2021

## Slide 9: Research Objective
✔️ Address direct ASC of PET images without CT anatomical data
✔️ Create a universal model applicable to different radiotracers and scanners ✔️ Evaluate model performance in artifact removal across a multi-center dataset

## Slide 10: Materials and Methods

## Slide 11: Workflow of A CT-less deep learning network
PET raw data
CT images
Recon without AC & SC
NAC (Input)
Model Training
DL PET
Recon with AC & SC
MAC (Target)
Recon : reconstruction;
AC : attenuation correction,
SC : scatter correction,
NAC : non-corrected PET,
MAC : CT-Measured attenuation/scatter corrected,
DL PET : deep learning–corrected PET,
Direct Method
[2] Guo, Nat Commun. 2022
Notes: A simple workflow of our work is : we used reconstructed pet images without applying any correction as input of a neural network. And we used a ct based corrected version of imput images as our target. From now on we named out imput nac and the target of the model or rather ground truth as mac.

## Slide 12: Workflow of A CT-less deep learning network
PET raw data
CT images
Recon without AC & SC
NAC (Input)
Model Training
DL PET
Recon with AC & SC
ADCM (Target)
Semi-Direct Method
Anatomy-Dependent Correction Maps
Radiotracer-Dependent textures
[2] Guo, Nat Commun. 2022
Notes: A simple workflow of our work is : we used reconstructed pet images without applying any correction as input of a neural network. And we used a ct based corrected version of imput images as our target. From now on we named out imput nac and the target of the model or rather ground truth as mac.

## Slide 13: Sample of Input and Ground truths of the network
NAC MAC ADCM

## Slide 14: 68Ga PET/CT Dataset [4]
Center | No | Train | Validation | Test | Scanner | Reconstruction | Matrix size × Z*
Center 1 | 56 | 43 | 11 | 2 | Siemens Biograph 6 | 3D-OSEM | 168 × 168
Center 2 | 31 | 25 | 4 | 2 | GE Discovery IQ | 3D-OSEM | 192 × 192
Center 3 | 45 | 35 | 8 | 2 | Siemens mCT | 3D-OSEM | 200 × 200
Center 4 | 40 | 28 | 10 | 2 | Siemens Biograph 6 | 3D-OSEM | 168 × 168
External Center | 12 | - | - | 12 | Siemens Horizon | PSF+TOF+3D-OSEM | 180 × 180
Total | 184 | 131 | 33 | 20 | - | - | -

## Slide 15: 18F-FDG Dataset [5]
Center | No | Train | Validation | Test | Scanner | Reconstruction | Matrix size × Z*
Center 6 | 55 | 39 | 6 | 11 | Siemens Biograph 6 | 3D-OSEM | 168 × 168
Center 7 | 43 | 23 | 9 | 10 | GE Discovery IQ | 3D-OSEM | 192 × 192
Total | 98 | 62 | 15 | 21 | - | - | -

## Slide 16: Artifact dataset
190 patients 18Ga PSMA PET/CT Mostly Motion and Halo artifacts

## Slide 17: Hyper parameters
Dyn-UNet architecture
3D patches sized 168×168×16
20 patches per patient
Loss function of the mean squared error (MSE)
Learning rate of 0.001
Optimization Adam algorithm
Batch size: 4
spacing : [4.07, 4.07, 3.00]
four layers: [16, 32, 64, 128]
Kernels: [3, 3, 3]
2 super vision head
The architecture of DynUnet

## Slide 18
IMCM  Integrated multi-Centre model ADCM  Anatomy-Dependent Correction Model TL-MC  Transfer Learning for IMCM model
68Ga-PET
IMCM
ADCM
18F-FDG
TL-MC
ADCM
NAC MAC IMCM TL - MC ADCM

## Slide 19: Quantification Metric

## Slide 20: Results

## Slide 21: 68Ga-PET

## Slide 22: 68Ga-PET

## Slide 23: 18F-FDG
Notes: For ME, ADCM shows overestimated and TLMC shows understination,
For RE, RMSE TLMC shows lower values.
About PSNR and SSI, ADCM shows more stability and constant parameters regarding TL-MC.

## Slide 24: 18F-FDG

## Slide 25
Case Study on artifact images

## Slide 26: Halo artifact in kidney

## Slide 27: Halo artifact

## Slide 28: Motion Artifact

## Slide 29: Motion & Halo Artifact

## Slide 30: Motion artifact

## Slide 31
Discussion

## Slide 32
🗨️ Developed a CT-free DL models for ASC in PET imaging,
🗨️ Tested across different centers and radiotracers.
🗨️ In Face of cross-center data, IMCM outperformed ADCM, in MEA, RE, RMSE and SSI.
Regarding cross-tracer data
Discussion
Notes: 🗨️ Developed a CT-free DL models for ASC in PET imaging, tested across different centers and radiotracers.
🗨️ IMCM outperformed ADCM, demonstrated more reliable and consistent performance in SUV estimations.
🗨️ IMCM showed significant improvement in quantitative accuracy and artifact correction.
🗨️ Using large, heterogeneous datasets enhances the generalizability and applicability of ASC models in PET imaging. These diverse datasets ensure the model performs well on unseen data, making it robust and reliable for real-world clinical settings.

## Slide 33
Limitations

## Slide 34
Limitations
🧱 Single universal models may not be effective; radioisotope-specific tuning is necessary.
🧱 Limited access to computational resources caused time consuming each training loop.
Notes: 🧱 Models optimized for one radiotracer type may not perform well with others, necessitating tracer-specific tuning.
🧱 Single universal models may not be effective; radioisotope-specific tuning is necessary.
🧱 Limited access to computational servers caused each training loop to take excessive time.

## Slide 35
Future Work

## Slide 36
📍 Train and test the model on more data from other centers.
📍 Assess the model organ-wise.
📍 Explore more Clinical Imaging Parameters.
📍 Targeted analysis for artifactual images.
📍 Rigorous statistical testing of categorized outcomes.
📍 Investigate other Neural Networks.
📍 Explore more other hyper-parameters.
Future Work
Notes: 📍 Train and test the model on more data from other centers.
📍 Assess the model organ-wise.
📍 Explore more Clinical Imaging Parameters: SUVmean, SUVmax, and total lesion metabolism.
📍 Targeted analysis for artifactual images.
📍 Rigorous statistical testing of categorized outcomes.
📍 Investigate Alternative Neural Networks generative adversarial networks (GANs), variational autoencoders, transformer-based models like Swin-UNet.
📍 Experiment with other optimizers like AdamW, RMSprop, and Nadam.

## Slide 37
Conclusion

## Slide 38
🔍 Designed a network to directly performs ASC from raw data.
This approach is less vulnerable to the challenges of pseudo-CT synthesis.
🔍 Preformed well in different centers.
making it reliable and flexible for use in clinics.
🔍 Addressed the artifact challenges.
high capability in detecting and removing artifacts in PET scans.
🔍 Constructed an assurance method
to not only detect but to also overcome those artifacts.
🔍 Be applicable where only Stand-alone PET
or PET/MRI systems are available.
🔍 Reduced in the cumulative radiation dose
would be valuable, particularly the paediatrics population
Conclusion
Notes: 🔍 Designed a network to directly performs ASC from raw data. This approach is less vulnerable to the challenges of pseudo-CT synthesis.
🔍 Preformed well in different centers, making it reliable and flexible for use in clinics.
🔍 Addressed the artifact challenges, high capability in detecting and removing artifacts in PET scans.
🔍 Constructed and validated a highly effective and fast quality assurance method to not only detect but to also overcome those artifacts.
🔍 Be applicable where only Stand-alone PET or PET/MRI systems are available.

## Slide 39: References
[4] Shiri I, Salimi Y, Hervier E, Pezzoni A, Sanaat A, Mostafaei S, et al. Artificial Intelligence-Driven Single-Shot PET Image Artifact Detection and Disentanglement: Toward Routine Clinical Image Quality Assurance. Clin Nucl Med. 2023 Dec 1;48(12):1035–46
[1] Schmitz et al. The physics of PET/Ct scanners. PET and PET/CT: a clinical guide, 2005. [2] Chen, X., Liu, C. Deep-learning-based methods of attenuation correction for SPECT and PET. J. Nucl. Cardiol. 30, 1859–1878 (2023). [3] Guo R, Xue S, Hu J, Sari H, Mingels C, Zeimpekis K, et al. Using domain knowledge for robust and generalizable deep learning-based CT-free PET attenuation and scatter correction. Nat Commun. 2022 Dec 1;13(1). [5] Shiri I, Salimi Y, Hervier E, Pezzoni A, Sanaat A, Mostafaei S, et al. Artificial Intelligence-Driven Single-Shot PET Image Artifact Detection and Disentanglement: Toward Routine Clinical Image Quality Assurance. Clin Nucl Med. 2023 Dec 1;48(12):1035–46.

## Slide 41: Workflow of Decomposition-Based Approach [1]
NAC (Input)
Model Training
DL PET
MAC (Target)
Anatomy-Dependent Correction Maps (ADCM)
Anatomy-independent textures
[2] Guo, Nat Commun. 2022

## Slide 42: Workflow of Decomposition-Based Approach [1]
[2] Guo, Nat Commun. 2022
NAC (Input)
ADCM
Model Training
DL PET
Anatomy-Dependent Correction Maps (ADCM)
Anatomy-independent textures

## Slide 43: Workflow of Decomposition-Based Approach [1]
NAC (Input)
ADCM
Model Training
DL ADCM
DL PET
[2] Guo, Nat Commun. 2022

## Slide 44
Integrated multi-Centre model (IMCM): the best metric of 0.0527 at epoch 434.
68Ga-PET

## Slide 45
Anatomy-Dependent Correction Model (ADCM): the best metric of 0.1237 was reached at epoch 466.
68Ga-PET

## Slide 46
Motion and Halo

## Slide 47
Transfer Learning for IMCM model (TL-MC):
18F-FDG

## Slide 48
Truncation and Motion

## Slide 49
Halo

## Slide 50
Motion Artifact

## Slide 51
Metal artifact

## Slide 52
Motion Artifact???

## Slide 53
???

## Slide 54
??